## Slide 1: Sterilization Monitoring
Pamela Alston, DDS
Lead Oral Health Specialist, Humanitas, Inc.
August 11, 2021

## Slide 2: To have your attendance recorded in the minutes
Attendance is taken by a survey.
The survey link is placed in the chat box during the teleconference
It is also read aloud at least once.
If an attendee is not able to go to the link, the attendee should email Leah Pan at: Leah.Pan@humanitas.com
Each attendee must complete a separate survey.
It is recommended that you complete the survey link within 48 hours after the teleconference.
Once the minutes are submitted to the National Office, there is no opportunity to make edits. Attendance cannot be updated

## Slide 3: Objectives
Review PRH requirements pertaining to monitoring sterilization
What to do in the event of a positive spore test

## Slide 4: R17. Communicable Disease and Infection Control
The center must:
b. Manage all cases of communicable disease and use protective measures as recommended by the Centers for Disease Control and Prevention (CDC).
c. Biologically monitor the function of autoclaves and maintain a log of spore test results.
d. Follow infection control measures as mandated by state and federal law.
Reference: PRH Chapter 2. Student Support Services, Section 2.3 Health Services

## Slide 5: CDC’s Guidelines for Infection Control in Dental Health Care Setting--2003
Based on unchanged scientific evidence
Represent minimum standards of practice
Designed to prevent infectious disease transmission
Non-regulatory but enforceable for JC

## Slide 6: Label instrument packages
Date of sterilization
Cycle of load number
Help in retrieving processed packages in event of a sterilization failure

## Slide 7: Spore Tests
Known as biological monitoring
Needs to be done at least weekly
When spores are killed during a sterilization cycle, it is assumed that all microorganisms have been destroyed and sterilization is achieved.
Downside—results not obtained immediately
Monitor the sterilization process by killing highly resistant microorganisms

## Slide 8: Mechanical monitoring
Doesn’t guarantee sterilization but helps detect procedural errors/equipment malfunctions
Involves checking the sterilizer gauges
Documents sterilization pressure, temperature, and exposure time
Failure to achieve any of those parameters is a first indication of a problem

## Slide 9: Chemical Monitoring
Sensitive chemicals that change color
e.g., indicator tapes, strips, tables and special markings
Results are immediate following the cycle
Provide more information than spore test
Indicator should be place in every package
Check for color change after cycle

## Slide 10: Single Parameter vs. Multiparameter Chemical Indicators
One sterilization parameter—time or temp
Multiparameter—time & temperature, presence of steam

## Slide 11: In Summary…
All 3 forms of sterilization monitoring are important
Document that sterilization cycle parameters have been met.
Keep spore test results. (no guidance)
Have them available for Program Compliance Assessments
Also check what your state dental bd. says

## Slide 12: Common factors for sterilization failures
Chamber overload
Excessive packaging material
Inadequate exposure time
Incorrect temperature and/or pressure settings
Failure to preheat the sterilizer
Interruption of the sterilization cycle

## Slide 13: “If the contaminated instruments were properly cleaned and the temperature was achieved as shown by the acceptable chemical (internal or external) indicator and the mechanical indicator that indicates time, temperature and pressure, a positive spore test will probably not indicate a sterilizer malfunction.” --OSHA Review

## Slide 14: “The margin of safety is sufficiently large that there is minimal infection risk associated with items in a load that show spore growth, especially if the item was properly cleaned and the temperature was achieved (e.g., as shown by acceptable chemical indicator or temperature chart). However, the spore test should be repeated immediately after correctly loading the sterilizer and using the same cycle that produced the failure.” --OSHA Review

## Slide 15: While the sterilizer is being retested…
Remove the sterilizer from service
Review all chemical and mechanical monitoring records since last – spore test
Review sterilization procedures with staff

## Slide 16: Common reasons for + spore test in absence of mechanical failure
Overloading the chamber
Failure to provide adequate package separation
Incorrect or excessive packaging
Put sterilizer back into service if the repeat spore test is negative

## Slide 17: Other steps in the case of a positive spore test
Get a loaner from the repair company
2nd test negative, put sterilizer back in service
2nd test positive, get sterilizer inspected/repaired by a repair company
Rechallenge with spore tests in 3 consecutive empty chamber sterilization cycles.
Recall and resterilize packages sterilized after last negative spore test